Dear Dave,

Please, find attached the final material on the infrared telemetry system. I hope this way of presentation will meet your requirements. If not, would you, please, specify, what kind of information you need.

Hope to hear from you soon.

Yours Sincerely,

Rita

Multifunctional infrared telemetry system for medical monitoring Working Prototype

Background and Significance

In the closed environments packed with electrical and electronic equipment the shortcomings of radio-telemetry become especially apparent, main of them being radio-wave interference and electromagnetic noises which significantly decrease reliability and accuracy of communication.

The multifunctional infrared telemetry system ensures a reliable wireless simultaneous transmission of several physiologic indices within the limited closed environment under the conditions of free movement and natural motor activity of the patients performing functional tests, professional tests or other diagnostic, preventive or treatment operations.

In environments with the high level of electromagnetic interference the main advantage of the infrared communication channel is its invulnerability to electromagnetic noises because of the considerable distance between the infrared and electromagnetic spectrums.

Innovative features

Utilization of the infrared communication channel and specific technical solutions enable the infrared telemetry system to simultaneously encode several (up to 6) physiological indices and transmit the encoded information via 6 channels using time-division of the channels, to pick up dispersed infrared radiation by the sensitive photo detectors.

The indices are encoded and transmitted using impulse regimen of infrared radiation and time-impulse way of encoding.

The system provides continuos monitoring of one subject for, at least, one hour.

General use

The device was developed to transmit the following indices: electrocardiogram, electroencephalogram, electromyogram, electro-oculogram, pulmonary ventilation indices (rate of respiration, volume of inspired/expired air), mechanical pulsation of the cardiovascular system, cutaneogalvanic reactions, kineto- and ballistocardiograms.

The device was conceived as an open system with the possibility for introducing additional indices and new methods of diagnosis, therefore, it can be adapted for particular requirements of the user.

At present, the following working units are completed:

transmitting unit worn by the subject:

·	pickup device,

·	encoding device

·	radiation device

stationary receiving unit:

·	photo detectors,

·	system for information transformation and visualization.

Brief technical characteristics

The present multifunctional infrared telemetry system is an analogue system.

Transmitting unit

Main characteristics of the preamplification channels:

·	maximal coefficient for voltage gain - 300,

·	input resistance - not less than 100 M,

·	the lowest boarder of the transmission band - 0,5 Hz  20%,

·	the upper boarder of the transmission band - 500 Hz  10%,

·	coefficient for cophasal interference suppression - not less than 75 dB.

The channel switching of the transmitting unit provides time division of the channels with a polling frequency not less than 1000 polls per second.

Receiving unit

Sensitivity of the receiving photo detector is about 6 mA/lm.

Output voltage of the information channels 0  6 V with a load of 1 k.

Power supply - 220 V, 50 Hz, 0,6 A.

Required equipment

Radio measuring equipment

PC Intel Pentium II, 266 MMX

Accessories and materials: serial electronic and electronic-radio elements, electronic-radio units, primary converters of the physiologic information (sensors, electrodes), other materials.

Funding

Salary	$5000

Equipment	$ 4000

Accessories	$ 1000

Required time

12 months

Expected Results

Utilization of the infrared communication channel invulnerable to electromagnetic noises ensures a high accuracy of information transmission, therefore the multifunctional infrared telemetry system will provide new quality and high efficiency of instrumental methods for evaluating functional state of the subjects under the conditions of required motor activity in specific environment.

The working prototype will provide simultaneous monitoring of 6 physiologic indices and would be used for making a production design. Production of the infrared telemetry system will find large market due to the multifunctional and wide range of the system applications.